Almost All Digital Electronics L/C Meter IIB
Neil Hecht (? to Aug 19, 2015)
Nuts & Volts April 1997

Build this self-calibrating, autoranging inductance/capacitance meter and sort out all those rf and audio chokes and capacitors you have stashed in your junk box. Its range -- from .001 UHY to 100 MHY and .01 PF to 1 UFD -- covers the vast majority of components for rf and audio filters, amplifiers, and oscillators.

Back in the 1950s, my dad and I built a lot of amateur radio equipment which involved winding our own coils for tuned circuits. The one thing we always wished for was a piece of test equipment to measure small values of inductance and capacitance. More than 40 years later, I have developed just such a piece of equipment, and -- although my dad is no longer with me -- I'm sure that other amateurs, engineers, and electronics enthusiasts will appreciate being able to measure small inductors and capacitors.

L/C Meter IIB is a handheld digital inductance/capacitance meter that features AUTOMATIC RANGING and SELF-CALIBRATION. If you do a good job of soldering, it is a "turn it on and it works" project.

Specifications
The unit displays values in one of two modes which can be changed during operation. The "micro mode" displays values in uHy, mHy, pF, and uF when applicable. In this mode, for example, 10.00 nano-Farads displays as .01000 micro-Farads and 1 nano-Henry displays as .001 micro-Hy. It is for old timers like me and is the way many parts are marked. The "nano mode" is for those more metrically inclined. Table 1 shows how each range is displayed in each mode.

One input pin to the micro-controller is a jumper which determines if the unit powers up in the "micro mode" (jumper open) or "nano mode" (jumper shorted).

Operating Modes
When the Lx and Cx switches are off, pressing the ZERO button sequences L/C
Meter IIB through five different operating modes:

READY MEASURE n measures Lx or Cx and displays the result in "nano mode"
(i.e. Lx = 99 nHy, Cx = 12.34 nF).

READY MEASURE u measures Lx or Cx and displays the result in "micro mode"
(i.e. Lx = .099 uHy, Cx = .01234 uF).

READY MATCHnMODE first measures a reference component Lz or Cz and displays
the value in "nano mode." When the ZERO button is pressed, this value is
stored in RAM and the difference between it and subsequent components is
displayed in "nano mode" (i.e., Lx-Lz = ??? nHy, Cx-Cz = 12.34 nF).

READY MATCHuMODE first measures a reference component Lz or Cz and displays
the value in "micro mode." When the ZERO button is pressed, this value is
stored in RAM and the difference between it and subsequent components is
displayed in "micro mode" (i.e., Lx-Lz = .099 uHy, Cx-Cz = .01234 uF).

READY MATCH%MODE first measures ta reference component Lz or Cz and displays
the value in "nano mode." When the ZERO button is pressed, this value is
stored in RAM and the ratio of the difference between it and subsequent
components is displayed in percent (i.e., (Lx-Lz)/Lz*100 = 12.34%,
(Cx-Cz)/Cz*100 = 12.34%).

Note that a positive reading in the matching modes means Lx is greater than
Lz, or Cx is greater than Cz, and vise versa.

The L/C Meter IIB is intended to measure inductors and capacitors "out of the
circuit." Inductors must have a reasonable Q for their value and negligible
distributed capacitance for their value. I have tested it using commercially available RF chokes ranging from 0.1 micro-Henry to 1000 micro-Henry. Hash chokes up to 100 micro-Henry wound on ferrite rods; on Pi-wound, RF chokes up to 7.5 milli-Henry; on toroid-wound inductors, up to 150 milli-Henry (such as the Hi-Q series obtainable from Mouser Electronics); and on several slug-tuned inductors from a Coilcraft Slot-10 designers kit (similar to the TOKO line of tunable inductors).

Circuit Description

The Oscillator
The key to L/C Meter IIB's operation is the oscillator circuit in Figure 1.
The LM311 is a voltage comparator. When power is applied, the voltage at pin
2 is 2.5 volts, causing the output to be at a level of five volts. This
charges capacitor C4 through resistor R4 until the voltage at pin 3 equals 2.5
volts.

As it reaches 2.5 volts, the output switches to a low level, inducing a
transient into the tank circuit composed of L1 and C1. The transient causes
the tuned circuit to ring at its resonant frequency. The ringing causes a
squarewave at the resonant frequency to appear at the output of the voltage
comparator. The squarewave is coupled back to the tuned circuit through R3
and C3 sustaining oscillation.

For the nominal values of L1 (68 uHy) and C1 (680 pF), an increase in L of 1
nHy (.001 uHy) or an increase in C of .01 pF produces a freqeuncy change of
slightly more than 5 Hz. A 0.2-second measuring period can resolve 5 Hz and
therefore .001 uHy or .01 pF.

Besides being simple, this oscillator circuit is very reliable in that it always starts and can tolerate a large variation in the inductance and capacitance used in the tank circuit.

The Micro-Computer
The complete schematic is shown in Figure 2. The output of the oscillator is applied to the RTCC (Real Time Clock Count) pin. This increments an eight-bit counter inside the microcomputer. The microcomputer accumulates the count for a period of 0.2 seconds. The frequency is then the accumulated count divided by the period. Discrete signals from the Lx, Cx, and ZERO switches are input to the microcomputer so knows what the operator wishes to do.

U2, a PIC16C622 microcomputer, is L/C Meter IIB's brain. The PIC16C622 is an
advanced version of the familiar PIC16C54 18-pin microcomputer. The 16C622
has a 14-bit instruction allowing CALLs and GOTOs to anywhere in its 2048
instruction program memory without the page management overhead of the 16C54.

It has 128 bytes of RAM and an eight-level deep stack rather than the
two-level stack of the 16C54. It also has interrupts which are not used in
L/C Meter IIB. Another useful feature is built-in pull-up resistors on the
inputs which helps reduce the parts count.

Self-Calibrating
During the calibrate cycle, the microcomputer first measures F1, the frequency
when only L1 and C1 are in the tank circuit. The frequency will be:
F1= _____1__________
    2 pi sqrt(L1 C1)

In order to obtain another equation, so that we can solve for both L1 and C1,
a known capacitor is switched into the tank circuit. The microcomputer
energizes relay RLY1 by raising the CALIB line to a logic high level which
switches capacitor C2 into the tank circuit. C2 is a .5% tolerance 1020 pF
capacitor composed of C2a and C2b. C2a is a 1000pF 2% polystyrene capacitor.
C2b is a selected NPO ceramic which, in parallel with C2a, equals 1020 pF
+-5 pF. This causes the frequency to become:
F2= ______1_______________
    2 pi sqrt(L1(C1+1020))

The two equations can be solved simultaneously to give:

C1= ______F2^2_____ 1020pF
    (F1^2 - F2^2)
and finally

L1 = _______1______
     4 pi^2 F1^2 C1

Because of the self-calibration capability, the exact values of L1 and C1 are
not critical, and 10% tolerance components are used. The accuracy of the
device is dependent upon C2, which is a .5% tolerance capacitor combination.

Measuring Inductance and Capacitance
When the Lx and Cx switches are off, the microcomputer continuously measures
F1 to track any drift in frequency. When the Lx switch is depressed, the
unknown inductor is placed in series with L1. The total inductance is then
L1 + Lx. This causes the frequency to change to:
F2 = _________1___________
     2 pi sqrt((L1 + Lx) C1)

This equation can be solved simultaneously with the equation for F1:
F1 = _______1________
     2 pi sqrt(L1 C1)

to produce:

     | F1^2 |
Lx = | ____ -1 | L1
     | F2^2 |

Similarly, when the Cx switch is depressed, the unknown capacitor is placed in
parallel with C1. The total capacitance is then C1 + Cx.
F2 = __________1_________
     w pi sqrt( L1(C1+Cx))

to produce:

     | F1^2 |
Cx = | ____ -1 | C1
     | F2^2 |

Stray Inductance and Capacitance
The circuit traces on the PCB, the switches, and the test leads all contribute
a small amount of "stray" inductance (Ls) and capacitance (Cs). These stray
add to the values of L1 or C1 when the Lx or Cx switches are pressed, slightly
effecting the frequency of F1.

The unit is zeroed by pressing the ZERO switch which causes the unit to store
the values of stray inductance or capacitance and subtracts them from the
measured values. The values displayed are:

     | F1^2 |
Lx = | ____ -1 | L1 - Ls
     | F2^2 |

and:

     | F1^2 |
Cx = | ____ -1 | C1 - Cs
     | F2^2 |

To zero Ls, the operator must short-circuit the test leads, press Lx and then
press the ZERO button. Similarly for capacitors, the operator open circuits
the test leads, presses Cx, and then presses ZERO.

The stored values of Ls and Cs are saved until the operating mode is changed.
When measuring coponents, it is not necessary to re-ZERO between components.
When the operating mode is changed from MEASURE to MATCH, these values are
reset to zero.

You will notice from the previous equations that inserting an unknown always
causes F2 to be less than F1. If an inductor is inserted when the Cx switch
is depressed, the result will be an increase in frequency -- F2 greater than
F1 --rather than a decrease. This is because the inductor has been placed in
parallel with L1 and inductors in parallel are always less than the value of
the smallest of the two values.

If the unit detects an increase in frequency, it will display "NOT A
CAPACITOR." This does not work for very large values of Lx. The decrease in
the effective value of L1 is trivial, while the shunt capacitance of the
large inductor is significant, and the frequency will decrease causing an
erroneous reading.

The effect of putting a capacitor in when the Lx switch is pressed is similar,
except the oscillator tends to stop, causing F2 to be zero. The unit detects
this and displays "NOT AN INDUCTOR." This is not true for very large values
of Cx in which case the unit may display an erroneous reading.

L/C Meter IIB can zer out ANY value in its range. If a value is inserted and
ZERO'd, the unit will display the difference between it and subsequent
components similar to the MATCHnMODE and MATCHuMODEs. The difference in the
MATCHxMODEs is that the range is frozen to the resolution of the initial
component. This limits the minimum difference in values to be one part in
10,000, or .01%. The reason for this may not be obvious. The maximum
resolution of the unit is four digits <i> at the value of the components
being measured. </i>

Consider two components: one with an exact value of 5000 pF, and the other
with an exact value of 5010.25 pF. The difference would be 10.25 pF; however,
the unit cannot resolve less than 1 pF at this range, and it would be
misleading to display the fractional portion of the difference.

Floating-Point Math
From all of the previous equations, it would seem that there is some relatively
high-powered math involved -- and there is. The upper portion of program
memory in the microcompute contains a complete 32-bit floating-point math
package. The math package includes ADD, SUBTRACT, MULTIPLY, and DIVIDE. It
also contains conversion from INTEGER to FLOATING-POINT, FLOATING-POINT to
INTEGER, and INTEGER to BCD.

The computer measures frequency by counting the number of oscillator cycles for
a period of 0.2 seconds. The result is an integer. This number is converted
to floating-point, and all calculations are done in floating-point. Whe the
values of Lx or Cx are finally computed, the answer is converted to an integer
and then to BCD for display.

Construction
These instructions generally apply to a kit purchased from ADE, but will be of
considerable interest to scratch builders.

The unit is indeed simple, and there is no particular order of assembly.
Refer to Figure 3, <i>Parts Layout,</i> for important information.

<b>NOTE: There is only 3/8 inch space under the display. Leave enough lead
length to tip taller parts at an angle so that the vertical dimension does
not exceed 3/8 inch. </b>

Solder P1 -- the male square post header -- at the top of the PCB. Install
the contrast control, R6, on the back of the PCB -- otherwise you will not be
able to adjust it with the display installed. Install the two 3/4-inch
spacers for the test jacks, as shown in Figure 4, <i>Mechanical Detail.</i>
This should complete the PCB assembly.

Solder J1 to the display unit. If the display module has 16 holes, install
the display connector in pins 1-14 of the 16 holes in the display (left 14
holes when viewed from the front). Of interest to scratch builders -- who may
have an intelligent 16 character display -- there are two types of these. The
first is organized as two lines of eight characters; that is, the 16
characters are not in contiguous RAM locations. It can be recognized as
having one large IC on the back of the unit. That is the type for which
L/C Meter IIB's microcontroller is programmed.

The second is organized as one line of 16 characters in contiguous RAM
locations. It has two ICs on the back. Remember to remove the protective
film from the display module before putting on the front cover.

Decide which type of capacitor display you prefer the unit to be in upon
power-up. If you prefer the nF method, solder a jumper wire as indicated on
the parts layout.

Pass the leads from the battery clip through one of the slots in the battery
box of the case and solder them to the appropriate pads of the PCB. Plug in
the display, turn the contrast control fully clockwise, and turn on the unit.
The unit will display "WAIT" for 10 seconds, followed by "CALIBRATING" for two
seconds, followed by "READY MEASURE x."

If so, you're up and running. Adjust the contrast control so the background is
just barely visible. Install the PCB in the bottom of the case using three #4
sheet metal screws. Install the top cover of the case, and install the binding
posts as shown in Figure 4. Test leads should not exceed four inches in length
with a banana plug at one end and an alligator clip at the other.

It may be necessary to "move" the edge of a hole or slot in the case. This is
easily done using sandpaper, file, or hobby knife. Before fitting the test
jacks or screws in the back of the case, fit the cover and squeeze the case
together while testing the switches for binding o the edge of the slot. "Move"
the edge as required.

This PCB is single-sided and has one jumper wire. The kit unit uses double-
sided, plated thru-hole boards and no jumpers are required.

Troubleshooting
It is very unlikely you will have any problems; however, if you purchased your
kit from AADE, they will try to fix it free, except for a $4.00 return postage
and handling fee.

If it did not work, remove the PCB and carefully inspect to see if you have
soldered everything that should be soldered and not soldered anything that
should not be (look for solder bridges). Bad soldering accounts for 99% of
units failing to work immediately. Here are some hints on where to look.

<b> Blank display: </b> Contrast control not adjusted correctly. Start with
it fully clockwise.

<b> Blank display: </b> Check five-volt power to CPU and display.

<b> Displays eight black squares:</b> CPU not communicating with display.
Check solder around CPU and display. CPU crystal not oscillating. Check with
oscilloscope, if possible.

<b> Displays WAIT, then CALIBRATING and sticks in CALIBRATING: </b> Oscillator
(LM311) is not oscillating. Check soldering around LM311, LM311 properly
installed, parts properly installed. C3 in backwards?

<b> Displays WAIT, then CALIBRATING and sticks in CALIBRATING:</b> ZERO button
stuck in or not soldered. Check continuity to ground from pin 13 of the CPU.

Operation
The typical stray inductance is .04 to .06 uHys, and the typical stray
capacitance is 5 to 7 pFs. When measuring inductors less than 5 uHys or
capacitance less than 50 pFs, it is advisable to ZERO the unit first. For
larger values, the strays are insignificant to the result. It is difficult
to retain a reading of 0.000 pFs because of the extreme sensitivity of the
unit. Your body capacitance influences the reading. Try ZEROing the
capacitance and then move your hands around the test leads without touching
them. YOu will find you can adjust the reading a few hundreths of a pF.

To measure inductance, place the unknown across the test leads and depress Lx.
 To measure capacitance, place the unknown acrsoss the test leads and press
 Cx.

The oscillator tends to drift a few Hertz during the first few minutes of
operation. When measuring very small values, the unit should be allowed to
warm up for about five minutes. With a resolution of 5 Hz, the thermal drift
will always occur, as evidences by a slowly drifting reading. The first
readings after pressing Lx or Cx are the most accurate.

Accuracy and Resolution
L/C Meter IIB is specified at 1% of reading. I have about 60 components
which I had measured on an HP 4275A L/C Meter. Measuring these components on
L/C Meter IIB found an average error of 0.23% for inductors and 0.24% for
capacitors. These values ranged from .1 uHy to 6.8 mHy and 2.7 pF to .068 uFd.
Thse measurements were for a single unit and could vary from unit to unit by
.5% as a function of the exact value of C2.

L/C Meter IIB has a four-digit resolution which, for small values of L and C,
are 1 nHy and .01 pF. YOu cannot accurately measure values this small. The
resolution greatly exceeds the accuracy. You can measure values as small as
.01 uHy and .1 pF with about 15% accuracy. You generally won't find components
this small.

For example, a piece of wire less than one inch long is .01 uHy. The
resolution is, however, relative and can be used for sorting a batch of similar
components as it truly does indicate which are slightly larger or smaller than
others. Also, for small values of indictance, the leads will contribute quite
a bit to the value. Measuring from the ends of the leads instead of next to
the body of the component can add up to .25 uHy.

For small values, the frequency of operaton (test frequency) is about 750 KHz
decreasing to about 60 KHz at .1 uFds or 10 mHys and about 20 KHz at 1 uFd
or 100 mHys. <b>NV</b>

<b> PARTS LIST</b>
__________________

R1, R2, R3	100K ohm 1/4 watt
R4	47K ohm 1/4 watt
R5	1000 ohm 1/4 watt
R6	10K ohm potentiometer	Digi-Key D4AA14-ND
C1	680 pF
C2	1000 pF 2%	Mouser 140-PF2A 102F
C4,C5	.1 uFd ceramic
C3	10 uFd/10V tantalum
C3,C8,C9	10 uFd/10V electrolytic
C6, C7	22 pF ceramic
X1	8.0 MHz crystal
L1	100 uHy	Mouser 434-1120-101L
U1	LM311N voltage comp
U2	PIC16C622 microcomputer
U3	78L05 voltage regulator
RLY1	SPST N.O. reed relay	Digi-Key HE206
DISP	LM-16151 or equiv	Digi-Key OP116
J1	14-pin square post socket	Digi-Key 929974-01-36
P1	14-pin square post plug	Digi-Key S1022-36
Lx, Cx, PWR	DPDT alternate action SW	Digi-Key EG1001
ZERO	DPDT momentary SW	Digi-Key EG1002
SW buttons	Digi-Key EG1098
Case	PacTec HP9VB
Test Jacks	5-way binding posts

The following are available from:
    Almost All Digital Electronics (Visa/MC)
      1412 Elm St., S.E.
  Auburn, WA 98092
    Voice: 206-351-9316 9am-9pm (usually available on weekends)
      FAX: 206-931-1940
    (NOTE: our are code is scheduled to change to 253 after Apr. '97)
      E-Mail: neil@aade.com
  http://www.aade.com

Programmed U2	$29.95
Complete Kit	$89.95 (includes machined case with panel decal)
Assembled Unit	$109.95
(INCLUDE $4.00 shipping and handling USA and Canada, $10.50 else-
where.) (Washington State residents add 8% sales tax.)